GLC14080307 附件19                                                                                                   FM1408-50A

| 参考试块和代表件相似性验证记录 Validation Record of Similarity for Reference Standards and Representative Parts |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 仪器名称/型号/编号Equipment Name/Model/Serial No.： (1) 自动C扫描系统/SCM12Xi/S680 |  |  |  |  |  |  |  |
| 零件所属工作包：(2)B777FTB Package to which the part belongs: |  |  |  |  |  |  |  |
| 验证项 Validated Items |  |  | 衰减值 Attenuation |  | 噪音水平 Noise Level |  | 声速 Velocity |
| 零件编号 Part Number | (3)177W3103-301 |  | (4)35dB |  | (5) 35dB |  | (6)3000mm/s |
| 参考试块编号 Reference Standards Number | (7)SACC-FG-ALCORE-001 |  | (8) 35dB |  | (9) 35dB |  | (10) 3000mm/s |
| 差异 Discrepancy |  |  | (11)0 |  | (12) 0 |  | (13)0 |
| 是否有特殊公差要求？ Any special tolerance requirements? |  |  | (14)6dB |  | (15)N/A |  | (16) 5% |
| 验证结果 Validated Result |  |  | (17)符合 |  | (18) 符合 |  | (19) 符合 |
| 说明：客户规范无明确要求时，衰减差异应在±6dB内，声速偏差小于5%，噪音水平得到超声III级人员的认可。 Comment: Not clearly specified by customer specification, the attenuation difference shall be within ±6dB, velocity deviation shall be less than 5%, and the noise level shall be approved by UT level III. |  |  |  |  |  |  |  |
| 验证人员 Validated by |  | (20) |  | 验证日期 Validation Date |  | (21) |  |

GLC14040110 附件21                                                         FM1404-34A

填写说明Instructions for Filling：

本表格中的内容应由批准的超声III级人员填写。

The contents in this form shall be filled in by the approved ultrasonic Level III personnel.

- ：填写使用的仪器名称、型号和编号。

Fill in the name, model and number of the instrument used.

- ：填写验证零件所属工作包，如B777FTB。

Fill in the work package to which the validation part belongs, such as B777FTB.

- ：填写验证的零件编号。

Fill in the verified part number.

- ：填写零件的衰减值。

Fill in the attenuation value of the part.

- ：填写零件的噪音水平。

Fill in the noise level of the part.

- ：如适用，填写零件的声速。

If applicable, fill in the sound velocity of the part.

- ：填写对应参考试块的编号。阶梯状试块，还需填写阶梯号。

Fill in the corresponding reference block number. Step number shall also be filled in for stepped test block.

- ：填写对应参考试块的衰减值。

Fill in the attenuation value of the corresponding reference block.

- ：填写对应参考试块的噪音水平。

Fill in the noise level of the corresponding reference block.

- ：如适用，填写对应参考试块的声速。

If applicable, fill in the sound velocity of the corresponding reference test block.

- ：填写零件与对应参考试块的衰减差异值。

Fill in the attenuation difference between the part and the corresponding reference block.

- ：填写零件与对应参考试块的噪音水平差异性。

Fill in the noise level difference between the part and the corresponding reference block.

- ：填写零件与对应参考试块的声速差异值。

Fill in the sound velocity difference between the part and the corresponding reference block.

- ：填写衰减差异值是否有特殊公差要求。若有，填写特殊要求值。

Fill in whether there are special tolerance requirements for attenuation difference value. If yes, fill in the value of special requirements.

GLC14040110 附件21                                                         FM1404-34A

- ：填写噪音水平差异性是否有特殊公差要求。若有，填写特殊要求。

Fill in whether there are special tolerance requirements for noise level difference. If yes, fill in the special requirements.

- ：填写声速差异值是否有特殊公差要求。若有，填写特殊要求值。

Fill in whether there are special tolerance requirements for the sound velocity difference value. If yes, fill in the value of special requirements.

- ：填写衰减差异值验证结果，符合或不符合。

Fill in the attenuation difference value verification results, which are consistent or inconsistent.

- ：填写噪音水平差异性验证结果，符合或不符合。

Fill in the noise level difference verification results, which are consistent or inconsistent.

- ：填写声速差异值验证结果，符合或不符合。

Fill in the verification result of sound velocity difference value, which is consistent or not.

- ：验证人员签字。

Signature of the verifier.

- ：填写验证日期。

Fill in the validation date.